## Sheet1
Hazard ID | Regulation | Clause | Part/System Description | 4.2 Characteristic/ Intended use | Risk Analysis | [7 empty] | 5. Risk Evaluation |  | 6. Risk control
[5 empty] | 4.3 Hazard | 4.4 Estimate risk(s) for hazardous situation | [8 empty] | 6.2 Option Analysis |  | 6.3 Implementation of risk control measures | 6.4. Residual risk evaluation |  |  |  | 6.5 Benefit | 6.6 Other Hazards
[6 empty] | Normal/ Fault Condition (N/F)? | Initiating Events and Circumstances/Causes (Potential Cause of Hazard/ Failure) (원인) | Foreseeable sequences of Events | Hazardous situation | Harm | Initial Risk Estimation
[11 empty] | Sev. | Prob. | Risk | Acceptable Risk?(Y/N) | Risk control Measurment (Description of Mitigation or Criteria) D: Inherent safety by design P: Protective measures in the medical device itself or in the manufacturing process I: Information for safety |  | Verification/ Validation | Sev. | Prob. | Residual Risk | Accptable Risk?(Y/N)
HA-01 | ISO 14971 | C4 | What materials or components are utilized in the medical device or are used with, or are in contact with, the medical device? | Ultrasound gel | Biocompatibility (Toxicity) | N | Low quality Gel | Using inappropriate gel to patient's skin. | Applying inappropriate gel to patient's skin during diagnosis | Skin trouble | 3 | 2 | 6 | N | D | Using material meeting with type-certificate | Test report of Ultrasound gel | 3 | 1 | 3 | Y | N/A | No new and increased risks
P
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 2. Getting started)
HA-02 | ISO 14971 | C4 | What materials or components are utilized in the medical device or are used with, or are in contact with, the medical device? | Probe lens | Biocompatibility (Toxicity) | N | Low quality lens | Using inappropriate lens material to patient's skin. | Applying inappropriate gel to patient's skin during diagnosis | Skin trouble | 2 | 1 | 2 | Y | D | Using material meeting with type-certificate | Certificate of Probe | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HA-03 | ISO 14971 | C5 | Is energy delivered to or extracted from the patient? | - Ultrasound energy is delivered to the patient. - Ti / Mi value are indicated on screen. - N/A | Mechanical energy (Ultrasound energy) | N | Too high-intensity of ultraound energy | Unintended ultrasound energy radiation during diagnosis due to S/W error | Radiating unintended accuracy of ultrasound energy on the skin surface | Diagnostic failure | 3 | 2 | 6 | N | D | Functions that set the accuracy of contollable component in the software | Software Validation report (HRD-01-13) | 3 | 1 | 3 | Y | N/A | No new and increased risks
[15 empty] | P | Perform acoustic power and index and system acceptance testing to ensure that energy level is lower than defined limit | Acoustic output Report (HRD-01-21)
I
HA-04 | ISO 14971 | C15 | Is the medical device intended to be mobile or portable? | Mobile ME Equipment | Mechanical energy (Moving parts) | N | Lack of Wheel-lock's performance. | Brake failure | Physical impact on person by the equipment | Physical injury | 2 | 2 | 4 | Y | D | Using lock-type to all casters | Design V&V(Q5-29-041) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HA-05 | ISO 14971 | C33 | Is the medical device susceptible to environmental influences? | Operating Limits Shipping/Storage Limits: System Operating Limits: Transducer | Mechanical energy (Ultrasound energy) | N | Inappropriate storage of the ME equipment | Failure of the system | Electric shock on operators or patients due to the destroyed insulation. | Elect. Shock | 3 | 2 | 6 | N | D | Using material meeting with related public certificate | Critical component list (No. 277190) | 3 | 1 | 3 | Y | N/A | No new and increased risks
P
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 8. Specification)
HID-01 | IEC 60601-1 | 4.3 | Essential Performance | This system transmits ultrasound energy into the body of patients using 2D, M Mode, color Doppler to obtain ultrasound images. It can be used in the examination of abdomen, gynecology, small organs, vascular, urology, prostate, cavity, surgery, cardiac and pediatric examination. Applicable particular standard is EN 60601-2-37:2008+A11:2011
HID-02 | IEC 60601-1 | 4.7 | SINGLE FAULT CONDITION for ME EQUIPMENT | Heating test 1. closing a breach 2. fan lock Leakage Current test 1. Condition provided in a measuring instrument | Thermal energy (High temperature) Electromagnetic energy (Line voltage, Leakage current) | F | 1. closing a breach 2. fan lock 3. Condition provided in a measuring instrument | Malfunction of ultrasound control device | 1. A wrong diagnosis on a patient 2. Damages on a patient from ultrasound energy | Diagnostic error Burn
HID-03 | IEC 60601-1 | 7.2.2 | Identification | the ME EQUIPMENT | Information Hazard (Labelling) | N | barely legible | not or barely legible | Diagnosis by the ME equipment without labels | Diagnostic error | 2 | 3 | 6 | N | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Performing a safety test: Durability Marking test | Safety Test Report (No. 277190)
I
HID-04 | IEC 60601-1 | 7.2.2 | Identification | the ME EQUIPMENT | Information Hazard (Labelling) | N | Low adhesive property | not or barely legible | Diagnosis by the ME equipment without labels | Diagnostic error | 2 | 3 | 6 | N | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Performing a safety test: Durability Marking test | Safety Test Report (No. 277190)
I
HID-05 | IEC 60601-1 | 7.2.2 | Identification | the ME EQUIPMENT | Information Hazard (Labelling) | N | inadequate labeling | Wrong attachment during assembly | Diagnosis by the ME equipment with inadequate labels | Diagnostic error | 2 | 3 | 6 | N | D | Control the process progree of assembly with Assembly Guidance | Assembly Guidance (HRD-08-00) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HID-06 | IEC 60601-1 | 7.2.2 | Identification | the BF applied part(probe) is detachable components of the ME EQUIPMENT | Electromagnetic energy (Line voltage, Leakage current) | N | Inappropriate labelling | Used with undesignated devices. | Intended use is failed, and damage on patients' can occur due to the failure. | Elect. Shock | 3 | 1 | 3 | Y | D |  |  | 3 | 1 | 3 | Y | N/A | No new and increased risks
P
I
HID-07 | IEC 60601-1 | 7.2.2 | Identification | PEMS in the ME EQUIPMENT | Information Hazard (Specification of service and maintenance) | N | Uncertain S/W version | Misidentification of S/W version | Old S/W version used by service personnel | Diagnostic error | 2 | 1 | 2 | Y | D | Functions for checking DB size & compaction DB | Software Validation report (HRD-01-13) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HID-08 | IEC 60601-1 | 7.2.13 | Physiological effects-safety signs and warning | Probe emitting Ultrasound energy | Mechanical energy (Ultrasound energy) | N | Inappropriate safety signs and warning | Diagnosis by wrong method | contact the probe on the skin surface for a long time | Burn | 3 | 2 | 6 | N | D |  |  | 3 | 1 | 3 | Y | N/A | No new and increased risks
[15 empty] | P | Perform acoustic power and index and system acceptance testing to ensure that energy level is lower than defined limit | Acoustic output Report (HRD-01-21)
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 2. Getting Started)
HID-09 | IEC 60601-1 | 7.3.7 | Supply terminals | power terminal inside the ME EQUIPMENT | Electromagnetic energy (Line voltage, Leakage current) | N | Wrong AC cable connection by mis-use | Wrong cable connection during manufacturing | damages on assembler due to the wrong operation | Elect. Shock | 3 | 2 | 6 | N | D | Designing connectors to be contacted only to relevant areas. | Design V&V(Q5-29-051) | 3 | 1 | 3 | Y | N/A | No new and increased risks
P
I
HID-10 | IEC 60601-1 | 7.3.7 | Supply terminals | power terminal inside the ME EQUIPMENT | Electromagnetic energy (Line voltage, Leakage current) | N | Wrong DC cable connection by mis-use | Wrong cable connection during manufacturing | damages on assembler from the wrong operation | Elect. Shock | 3 | 1 | 3 | Y | D |  |  | 3 | 1 | 3 | Y | N/A | No new and increased risks
P
I
HID-11 | IEC 60601-1 | 7.4.2 | Control devices | Ultrasound frequency, power and etc. can be controllable | Mechanical energy (Ultrasound energy) | N | Excessive ultrasound energy radiation | Malfunction of ultrasound control device | Radiating excessive ultrasound energy on the skin surface | Burn | 4 | 3 | 12 | N | D | Functions that set the limit of contollable component in the software | Software Validation report (HRD-01-13) | 4 | 1 | 4 | Y | N/A | No new and increased risks
[15 empty] | P | Perform acoustic power and index and system acceptance testing to ensure that energy level is lower than defined limit | Acoustic output Report (HRD-01-21)
I
HID-12 | IEC 60601-1 | 7.9.3.2 | Replacement of fuses, POWER SUPPLY CORDS , other parts | Fuses and a power supply cord can be replacable | Electromagnetic energy (Line voltage, Leakage current) | N | Rplaceable parts | Replacement of fuses | Replace the fuses when operating | Elect. Shock | 3 | 3 | 9 | N | D | Use of medical grade AC Inlet | Critical component list (No. 277190) | 3 | 1 | 3 | Y | N/A | No new and increased risks
P
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 10. Safety)
HID-13 | IEC 60601-1 | 7.9.3.2 | Replacement of fuses, POWER SUPPLY CORDS , other parts | Fuses and a power supply cord can be replacable | Electromagnetic energy (Line voltage, Leakage current) | N | Rplaceable parts | Replacement of a power supply cord | Replace the a power supply cord when operating | Elect. Shock | 3 | 3 | 9 | N | D | non detachable power cord with fixed Guard. | Designed not to detach the fixed guard without tools Design V&V(Q5-29-045) | 3 | 1 | 3 | Y | N/A | No new and increased risks
P
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 10. Safety)
HID-14 | IEC 60601-1 | 8.8.4.1 | MECHANICAL strength and resistance to heat | External plastic covers | Electromagnetic energy (Line voltage, Leakage current) | N | Hazard voltage exposed due to faulty insulations | Insulations are destroyed from damages during their life span. | Electric shock on operators or patients due to the destroyed insulation. | Elect. Shock | 3 | 4 | 12 | N | D | Using insulated material meeting with related UL certificate | Critical component list (No. 277190) | 3 | 1 | 3 | Y | N/A | No new and increased risks
[15 empty] | P | Performing a safety test: 1. resistance to moisture 2. dielectric strength 3. mechanical strength | Safety Test Report (No. 277190)
I
HID-15 | IEC 60601-1 | 9.6.1 | Acoustic energy - General | Speakers | Mechanical energy (Acoustic energy) | N | Audible Noise Vibration | Unlimited sound from speaker | too much loud sound from the speaker during diagnosis | Damage to a ear of user and patient | 3 | 2 | 6 | N | D | Avoid vibration by using protective buffer between speaker and cover | Design V&V(Q5-29-042) | 3 | 1 | 3 | Y | N/A | No new and increased risks
P
I
HID-16 | IEC 60601-1 | 9.6.1 | Acoustic energy - General | Fan | Mechanical energy (Acoustic energy) | N | Audible Noise Vibration | Unlimited noise from fan | too much loud sound from the fan during diagnosis | Damage to a ear of user and patient | 2 | 1 | 2 | Y | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HID-17 | IEC 60601-1 | 9.6.2.2 | Infrasound and Ultrasound energy | Probe emitting Ultrasound energy | Mechanical energy (Ultrasound energy) | N | Excessive ultrasound energy radiation | Malfunction of ultrasound control device | Radiating excessive ultrasound energy on the skin surface | Burn | 4 | 3 | 12 | N | D | 1. Setting the limit of contollable component in the software 2. Designing the circuit with fixed voltage circuit to emit ultrasound | Software Validation report (HRD-01-13) Design V&V(Q5-29-038) | 4 | 1 | 4 | Y | N/A | No new and increased risks
[15 empty] | P | Perform acoustic power and index and system acceptance testing to ensure that energy level is lower than defined limit | Acoustic output Report (HRD-01-21)
I
HID-18 | IEC 60601-1 | 9.8.1 | MECHANICAL HAZARDS associated with support systems-General | Arm is supporting the monitor | Mechanical energy (Moving parts) | F | Supporting part failure | Monitor fall down due to the failure of arm when moving | Impact on user and patient's body during diagnosis | Physical injury | 4 | 2 | 8 | N | D | Designing monitor arm meeting with TENSILE SAFETY FACTOR | Design V&V(Q5-29-042) | 4 | 1 | 4 | Y | N/A | No new and increased risks
[15 empty] | P | Performing a safety test: 1. resistance to moisture 2. dielectric strength 3. mechanical strength | Safety Test Report (No. 277190)
I
HID-19 | IEC 60601-1 | 9.8.1 | MECHANICAL HAZARDS associated with support systems-General | Arm is supporting the monitor | Mechanical energy (Moving parts) | F | Unlimited movement radius of arm | Finger is pinched between upper arm and lower arm when diagnosis | Impact on user and patient's body during diagnosis | Physical injury | 2 | 3 | 6 | N | D | Designing monitor arm with locking system | Design V&V(Q5-29-042) | 2 | 2 | 4 | Y | N/A | No new and increased risks
[15 empty] | P | Performing a safety test: 1. resistance to moisture 2. dielectric strength 3. mechanical strength | Safety Test Report (No. 277190)
[15 empty] | I | Designating parts on the User Manual | User Manual (HM-01-0 Chapter 10. Safety) Label drawing (HRD-01-06(13))
HID-20 | IEC 60601-1 | 9.8.1 | MECHANICAL HAZARDS associated with support systems-General | Arm is supporting the monitor | Mechanical energy (Moving parts) | F | Unlimited turning radius of arm | Monitor turn around unlimitely due to the faulure of arm when moving | Impact on user and patient's body during moving | Physical injury | 4 | 2 | 8 | N | D | Designing monitor arm with locking system | Design V&V(Q5-29-042) | 4 | 1 | 4 | Y | N/A | No new and increased risks
[15 empty] | P | Performing a safety test: 1. resistance to moisture 2. dielectric strength 3. mechanical strength | Safety Test Report (No. 277190)
I
HID-21 | IEC 60601-1 | 9.8.1 | MECHANICAL HAZARDS associated with support systems-General | Arm is supporting the monitor | Mechanical energy (Moving parts) | F | Unlimited turning radius of arm | Monitor turn around unlimitely due to the faulure of arm when diagnosis | Impact on user and patient's body during diagnosis | Physical injury | 4 | 2 | 8 | N | D | Designing monitor arm with locking system | Design V&V(Q5-29-042) | 4 | 1 | 4 | Y | N/A | No new and increased risks
[15 empty] | P | Performing a safety test: 1. resistance to moisture 2. dielectric strength 3. mechanical strength | Safety Test Report (No. 277190)
I
HID-22 | IEC 60601-1 | 9.8.1 | MECHANICAL HAZARDS associated with support systems-General | Arm is supporting the monitor | Mechanical energy (Moving parts) | N | Complicated specification of assembly of arm | Wrong assembly by assembler on manufacture | Impact on user and patient's body during moving and diagnosis | Physical injury | 4 | 1 | 4 | Y | D |  |  | 4 | 1 | 4 | Y | N/A | No new and increased risks
P
I
HID-23 | IEC 60601-1 | 11.1.3 | Measurements | Probe can be contactable to patient's skin surface | Mechanical energy (Ultrasound energy) | N | Inappropriate diagnosis due to use error | Diagnosis by wrong method | contact the probe on the skin surface for a long time | Burn | 3 | 2 | 6 | N | D | Designing the circuit with fixed voltage to emit ultrasound | Design V&V(Q5-29-034) | 3 | 1 | 3 | Y | N/A | No new and increased risks
[15 empty] | P | Perform acoustic power and index and system acceptance testing to ensure that energy level is lower than defined limit | Acoustic output Report (HRD-01-21)
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 10. Safety)
HID-24 | IEC 60601-1 | 11.3 | Constructional requirements for fire ENCLOSURE of ME EQUIPMENT | Use the enclosure that meets requirements for fire | Biocompatibility (Toxicity) | F | Poor resistance to flame | Fire accident in hospital which has the ME equipment | Poisonous gas from the enclosure | Poison | 4 | 2 | 8 | N | D | Using material meeting with related UL certificate | Critical component list (No. 277190) | 4 | 1 | 4 | Y | N/A | No new and increased risks
P
I
HID-25 | IEC 60601-1 | 11.3 | Constructional requirements for fire ENCLOSURE of ME EQUIPMENT | Use the enclosure that meets requirements for fire | Thermal energy (High temperature) | F | Poor resistance to flame | Fire accident in hospital which has the ME equipment | Secondary fire from external conflagration | Fire | 4 | 2 | 8 | N | D | Using material meeting with related UL certificate | Critical component list (No. 277190) | 4 | 1 | 4 | Y | N/A | No new and increased risks
P
I
HID-26 | IEC 60601-1 | 11.6.8 | Compatibility with substances used with ME EQUIPMENT | Ultrasound gel | Electromagnetic energy (Line voltage, Leakage current) | N | Low quality Gel | Using inappropriate gel to probe. | Insulation breakdown | Elect. Shock | 2 | 1 | 2 | Y | D | Using material meeting with type-certificate | Test report of Ultrasound gel | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 2. Getting started)
HID-27 | IEC 60601-1 | 11.6.8 | Compatibility with substances used with ME EQUIPMENT | Ultrasound gel | Biocompatibility (Toxicity) | N | Low quality Gel | Using inappropriate gel to patient's skin. | Applying inappropriate gel to patient's skin during diagnosis | Skin trouble | 3 | 2 | 6 | N | D | Using material meeting with type-certificate | Test report of Ultrasound gel | 3 | 1 | 3 | Y | N/A | No new and increased risks
P
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 2. Getting started)
HID-28 | IEC 60601-1 | 11.6.8 | Compatibility with substances used with ME EQUIPMENT | Ultrasound gel | Use error (Knowledge-based failure) | N | Low quality Gel | Using inappropriate gel to patient's skin. | Wrong diagnosis due to poor transmission of ultrasound | Diagnostic failure | 3 | 2 | 6 | N | D | Using material meeting with type-certificate | Test report of Ultrasound gel | 3 | 1 | 3 | Y | N/A | No new and increased risks
P
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 2. Getting started)
HID-29 | IEC 60601-1 | 12.1 | Accuracy of controls and instrument | Ultrasound frequency, power and etc. can be controllable | Mechanical energy (Ultrasound energy) | N | Equipment faults, production faults, ageing | Malfunction of ultrasound control device | Radiating excessive ultrasound energy on the skin surface | Burn | 4 | 3 | 12 | N | D | Designing the circuit with fixed voltage to emit ultrasound | Design V&V(Q5-29-034) | 4 | 1 | 4 | Y | N/A | No new and increased risks
[15 empty] | P | Perform acoustic power and index and system acceptance testing to ensure that energy level is lower than defined limit | Acoustic output Report (HRD-01-21)
I
HID-30 | IEC 60601-1 | 12.1 | Accuracy of controls and instrument | Ultrasound frequency, power and etc. can be controllable | Mechanical energy (Ultrasound energy) | N | Equipment faults, production faults, ageing | Malfunction of ultrasound control device | Radiating unintended accuracy of ultrasound energy on the skin surface | Diagnostic failure | 3 | 3 | 9 | N | D | Functions that set the accuracy of contollable component in the software | Software Validation report (HRD-01-13) | 3 | 1 | 3 | Y | N/A | No new and increased risks
[15 empty] | P | Perform acoustic power and index and system acceptance testing to ensure that energy level is lower than defined limit | Acoustic output Report (HRD-01-21)
I
HID-31 | IEC 60601-1 | 12.4.2 | Indication relevant to safety | Ultrasound frequency, power and etc. are listed on the screen | Mechanical energy (Ultrasound energy) | N | Absence of indication relevant to Ultrasound power | Unintended excessive ultrasound energy radiation by user due to S/W error | Radiating excessive ultrasound energy on the skin surface | Burn | 4 | 3 | 12 | N | D | Functions that set and indicate the ultrasound power component in the software | Software Validation report (HRD-01-13) | 4 | 1 | 4 | Y | N/A | No new and increased risks
[15 empty] | P | Perform acoustic power and index and system acceptance testing to ensure that energy level is lower than defined limit | Acoustic output Report (HRD-01-21)
I
HID-32 | IEC 60601-1 | 12.4.2 | Indication relevant to safety | Ultrasound frequency, power and etc. are listed on the screen | Mechanical energy (Ultrasound energy) | N | Absence of indication relevant to Ultrasound frequency | Unintended excessive ultrasound energy radiation by user due to S/W error | Radiating excessive ultrasound energy on the skin surface | Burn | 4 | 3 | 12 | N | D | Functions that set and indicate the ultrasound frequency component in the software | Software Validation report (HRD-01-13) | 4 | 1 | 4 | Y | N/A | No new and increased risks
[15 empty] | P | Perform acoustic power and index and system acceptance testing to ensure that energy level is lower than defined limit | Acoustic output Report (HRD-01-21)
I
HID-33 | IEC 60601-1 | 12.4.2 | Indication relevant to safety | Ultrasound frequency, power and etc. are listed on the screen | Mechanical energy (Ultrasound energy) | N | Absence of indication relevant to MI, TI index | Unintended excessive ultrasound energy radiation by user due to S/W error | Radiating excessive ultrasound energy on the skin surface | Burn | 4 | 3 | 12 | N | D | Functions that set and indicate the MI, TI in the software | Software Validation report (HRD-01-13) | 4 | 1 | 4 | Y | N/A | No new and increased risks
[15 empty] | P | Perform acoustic power and index and system acceptance testing to ensure that energy level is lower than defined limit | Acoustic output Report (HRD-01-21)
I
HID-34 | IEC 60601-1 | 12.4.3 | Accidental selection of excessive output values | Ultrasound frequency, power and etc. can be controllable | Mechanical energy (Ultrasound energy) | N | Too low-intensity of ultraound energy | Unintended ultrasound energy radiation during diagnosis due to S/W error | Radiating unintended accuracy of ultrasound energy on the skin surface | Diagnostic failure | 3 | 2 | 6 | N | D | Functions that set the accuracy of contollable component in the software | Software Validation report (HRD-01-13) | 3 | 1 | 3 | Y | N/A | No new and increased risks
[15 empty] | P | Perform acoustic power and index and system acceptance testing to ensure that energy level is lower than defined limit | Acoustic output Report (HRD-01-21)
I
HID-35 | IEC 60601-1 | 12.4.3 | Accidental selection of excessive output values | Ultrasound frequency, power and etc. can be controllable | Mechanical energy (Ultrasound energy) | N | Too high-intensity of ultraound energy | Unintended ultrasound energy radiation during diagnosis due to S/W error | Radiating unintended accuracy of ultrasound energy on the skin surface | Diagnostic failure | 3 | 2 | 6 | N | D | Functions that set the accuracy of contollable component in the software | Software Validation report (HRD-01-13) | 3 | 1 | 3 | Y | N/A | No new and increased risks
[15 empty] | P | Perform acoustic power and index and system acceptance testing to ensure that energy level is lower than defined limit | Acoustic output Report (HRD-01-21)
I
HID-36 | IEC 60601-1 | 12.4.4 | Incorrect output | Probe emitting Ultrasound energy | Mechanical energy (Ultrasound energy) | N | Too low-intensity ultraound energy | Unintended ultrasound energy radiation during diagnosis due to S/W error | Radiating unintended accuracy of ultrasound energy on the skin surface | Diagnostic error | 1 | 2 | 2 | Y | D |  |  | 1 | 2 | 2 | Y | N/A | No new and increased risks
P
I
HID-37 | IEC 60601-1 | 12.4.4 | Incorrect output | Probe emitting Ultrasound energy | Mechanical energy (Ultrasound energy) | N | Too high-intensity ultraound energy | Unintended ultrasound energy radiation during diagnosis due to S/W error | Radiating excessive ultrasound energy on the skin surface | Burn | 4 | 3 | 12 | N | D | 1. Functions that set the accuracy of contollable component in the software 2. To alarm incorrect output to user, display MI, TI index on the screen | Software Validation report (HRD-01-13) | 4 | 1 | 4 | Y | N/A | No new and increased risks
[15 empty] | P | Perform acoustic power and index and system acceptance testing to ensure that energy level is lower than defined limit | Acoustic output Report (HRD-01-21)
I
HID-38 | IEC 60601-1 | 14 | PEMS in the ME equipment | MI (Mechanical Index) | Thermal energy (High temperature) | N | Unidentified MI Index | Repetition of the diagnosis due to the unidentified result | Radiating unintended ultrasound energy to the patient | Inconvenience during diagnosis | 2 | 3 | 6 | N | D | 1. Functions that set the accuracy of contollable component in the software 2. To alarm incorrect output to user, display MI, TI index on the screen | Software Validation report (HRD-01-13) | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Perform acoustic power and index and system acceptance testing to ensure that energy level is lower than defined limit | Acoustic output Report (HRD-01-21)
I
HID-39 | IEC 60601-1 | 14 | PEMS in the ME equipment | TI (Thermal Index) | Thermal energy (High temperature) | N | Unidentified TI Index | Repetition of the diagnosis due to the unidentified result | Radiating unintended ultrasound energy to the patient | Inconvenience during diagnosis | 2 | 3 | 6 | N | D | 1. Functions that set the accuracy of contollable component in the software 2. To alarm incorrect output to user, display MI, TI index on the screen | Software Validation report (HRD-01-13) | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Perform acoustic power and index and system acceptance testing to ensure that energy level is lower than defined limit | Acoustic output Report (HRD-01-21)
I
HID-40 | IEC 60601-1 | 14 | PEMS in the ME equipment | Power | Mechanical energy (Ultrasound energy) | N | Unidentified ultraound energy | Repetition of the diagnosis due to the unidentified result | Radiating unintended ultrasound energy to the patient | Inconvenience during diagnosis | 2 | 3 | 6 | N | D | 1. Functions that set the accuracy of contollable component in the software 2. To alarm incorrect output to user, display MI, TI index on the screen | Software Validation report (HRD-01-13) | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Perform acoustic power and index and system acceptance testing to ensure that energy level is lower than defined limit | 2. Acoustic output Report (HRD-01-21)
I
HID-41 | IEC 60601-1 | 14 | PEMS in the ME equipment | MI (Mechanical Index) | Function (Loss or deterioration of function) | N | Failure of the MI control Function | Unintended ultrasound energy radiation during diagnosis due to S/W error | Radiating excessive ultrasound energy on the skin surface | Diagnostic error | 2 | 2 | 4 | Y | D | Functions that set the accuracy of contollable component in the software | Software Validation report (HRD-01-13) | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Perform acoustic power and index and system acceptance testing to ensure that energy level is lower than defined limit | Acoustic output Report (HRD-01-21)
I
HID-42 | IEC 60601-1 | 14 | PEMS in the ME equipment | TI (Thermal Index) | Function (Loss or deterioration of function) | N | Failure of the TI control Function | Unintended ultrasound energy radiation during diagnosis due to S/W error | Radiating excessive ultrasound energy on the skin surface | Diagnostic error | 2 | 2 | 4 | Y | D | Functions that set the accuracy of contollable component in the software | Software Validation report (HRD-01-13) | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Perform acoustic power and index and system acceptance testing to ensure that energy level is lower than defined limit | Acoustic output Report (HRD-01-21)
I
HID-43 | IEC 60601-1 | 14 | PEMS in the ME equipment | Control Panel | Function (Loss or deterioration of function) | N | Loss of the Control Function | Failure of input from the user | Repetition of the diagnosis | Diagnostic error | 2 | 2 | 4 | Y | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (HRD-01-13, Chaptor 6)
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 2. Getting started)
HID-44 | IEC 60601-1 | 14 | PEMS in the ME equipment | RF Data | Function (Loss or deterioration of function) | N | Loss of the RF Data Function | Failure of input from the user | Repetition of the diagnosis | Diagnostic error | 2 | 3 | 6 | N | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (HRD-01-13, Chaptor 6)
I
HID-45 | IEC 60601-1 | 14 | PEMS in the ME equipment | Channel Data | Function (Loss or deterioration of function) | N | Loss of the Channel Datal Function | Failure of input from the user | Repetition of the diagnosis | Diagnostic error | 2 | 3 | 6 | N | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (HRD-01-13, Chaptor 6)
I
HID-46 | IEC 60601-1 | 14 | PEMS in the ME equipment | Trackball | Function (Loss or deterioration of function) | N | Loss of the Trackball Functionality Function | Failure of input from the user | Repetition of the diagnosis | Diagnostic error | 2 | 2 | 4 | Y | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (HRD-01-13, Chaptor 6)
I
HID-47 | IEC 60601-1 | 14 | PEMS in the ME equipment | Monitor | Function (Loss or deterioration of function) | N | Loss of the Monitor Functionality Function | Failure of input from the user | Repetition of the diagnosis | Diagnostic error | 2 | 2 | 4 | Y | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (HRD-01-13, Chaptor 6)
I
HID-48 | IEC 60601-1 | 14 | PEMS in the ME equipment | Printer | Function (Loss or deterioration of function) | N | Loss of the Printer Functionality Function | Failure of input from the user | Repetition of the diagnosis | Diagnostic error | 2 | 2 | 4 | Y | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (HRD-01-13, Chaptor 6)
I
HID-49 | IEC 60601-1 | 14 | PEMS in the ME equipment | Version management | Information Hazard (Specification of service and maintenance) | N | Uncertain S/W version | Misidentification of S/W version | Old S/W version used by service personnel | Diagnostic error | 2 | 1 | 2 | Y | D | Functions for checking DB size & compaction DB | Software Validation report (HRD-01-13) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HID-50 | IEC 60601-1 | 15.4.2.1 d | Loss of function of me equipment | There are fuses in AC Inlet and SMPS | Electromagnetic energy (Line voltage, Leakage current) | F | Poor choice of fuses in AC inlet which are not compatible to the ME Equipment | Fuses in AC inlet are failure | Absence of prevention from over current and voltage | Elect. Shock 1) Serious burns 2) Heart fibrillation 3) Death | 3 | 2 | 6 | N | D | Use of certified Equipment | Critical component list (No. 277190) | 3 | 1 | 3 | Y | N/A | No new and increased risks
P
I
HID-51 | IEC 60601-1 | 15.4.2.1 d | Loss of function of me equipment | There are fuses in AC Inlet and SMPS | Electromagnetic energy (Line voltage, Leakage current) | F | Poor choice of fuses in SMPS which is not compatible to the ME Equipment | Fuses in SMPS are failure | Absence of prevention from over current and voltage | Elect. Shock 1) Serious burns 2) Heart fibrillation 3) Death | 3 | 2 | 6 | N | D | Use of certified Equipment | Critical component list (No. 277190) | 3 | 1 | 3 | Y | N/A | No new and increased risks
P
I
HID-52 | IEC 60601-1 | 15.4.5 | Pre-set controls | Organ pre-set controls | Use error (Routine violation) | N | Patients of particular case who are not compatible to pre-set list | Diagnosis by wrong method | Diagnosis failure but even rare circumstances | Diagnostic error | 2 | 1 | 2 | Y | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HID-53 | IEC 60601-1 | 17 | Electromagnetic compatibility of ME equipment and ME systems | Applicable | Electromagnetic energy (Line voltage, Leakage current) | N | Electromagnetic not compatible | Electromagnetic interference from other equipment | Lose functions of the ME equipment in hospital environment which user uses other equipment together in | Diagnostic error | 3 | 2 | 6 | N | D |  |  | 3 | 1 | 3 | Y | N/A | No new and increased risks
[15 empty] | P | Using Shielded cable and ferrite core, comply with the requirementsas specified in IEC/EN60601-1-2 Group1, Class A device | EMC Test Report (NK-15-E-0048)
I
HID-54 | IEC 60601-1 | 17 | Electromagnetic compatibility of ME equipment and ME systems | Applicable | Electromagnetic energy (Line voltage, Leakage current) | F | Electrical Interference | 1. Influx of electromagnetic waves 2. Excessive Radiated Emission | Lose functions of the other ME equipment in hospital environment which user uses the ME equipment together in | Diagnostic error | 3 | 2 | 6 | N | D |  |  | 3 | 1 | 3 | Y | N/A | No new and increased risks
[15 empty] | P | Using Shielded cable and ferrite core, comply with the requirementsas specified in IEC/EN60601-1-2 Group1, Class A device | EMC Test Report (NK-15-E-0048)
I
HID-55 | IEC 60601-1 | 17 | Electromagnetic compatibility of ME equipment and ME systems | Applicable | Electromagnetic energy (Line voltage, Leakage current) | N | Magnetic Interference | influx of electromagnetic waves | Lose functions of the other ME equipment in hospital environment which user uses the ME equipment together in | Diagnostic error | 3 | 2 | 6 | N | D |  |  | 3 | 1 | 3 | Y | N/A | No new and increased risks
[15 empty] | P | Using Shielded cable and ferrite core, comply with the requirementsas specified in IEC/EN60601-1-2 Group1, Class A device | EMC Test Report (NK-15-E-0048)
I
HID-56 | IEC 60601-1 | 17 | Electromagnetic compatibility of ME equipment and ME systems | Applicable | Electromagnetic energy (Line voltage, Leakage current) | N | ESD (Elecro-Static discharge) due to difference of potential | Poor enclosure material near the monitor | The screen flickers due to absence of protective circuit | Diagnostic error | 3 | 2 | 6 | N | D | Use of certified Equipment | Critical component list (No. 277190) | 3 | 1 | 3 | Y | N/A | No new and increased risks
[15 empty] | P | Attaching protective tapes between monitor enclosures | EMC Test Report (NK-15-E-0048)
I
HID-57 | IEC 60601-1 | 17 | Electromagnetic compatibility of ME equipment and ME systems | Applicable | Electromagnetic energy (Line voltage, Leakage current) | N | Electro-magnetic immunity | influx of electromagnetic waves | Lose functions of the ME equipment in hospital environment which user uses other equipment together in | Diagnostic error | 3 | 2 | 6 | N | D |  |  | 3 | 1 | 3 | Y | N/A | No new and increased risks
[15 empty] | P | Using Shielded cable and ferrite core, comply with the requirements as specified in IEC/EN60601-1-2 Group1, Class A device | EMC Test Report (NK-15-E-0048)
I
HID-58 | IEC 60601-1 | 17 | Electromagnetic compatibility of ME equipment and ME systems | Applicable | Electromagnetic energy (Line voltage, Leakage current) | N | Electro-magnetic immunity | Excessive Radiated Emission | Lose functions of the ME equipment in hospital environment which user uses other equipment together in | Diagnostic error | 3 | 2 | 6 | N | D |  |  | 3 | 1 | 3 | Y | N/A | No new and increased risks
[15 empty] | P | Using Shielded cable and ferrite core, comply with the requirementsas specified in IEC/EN60601-1-2 Group1, Class A device | EMC Test Report (NK-15-E-0048)
I
HID-59 | IEC 60601-2-37 | 201.10.101 | Protection against unwanted and excessive radiation hazards | Probes | Mechanical energy (Ultrasound energy) | N | Too low-intensity of ultraound energy | Unintended ultrasound energy radiation during diagnosis due to S/W error | Radiating excessive ultrasound energy on the skin surface | Cell destruction | 3 | 2 | 6 | N | D |  |  | 3 | 1 | 3 | Y | N/A | No new and increased risks
[15 empty] | P | Perform acoustic power and index and system acceptance testing to ensure that energy level is lower than defined limit | Acoustic output Report (HRD-01-21)
I
HID-60 | IEC 60601-2-37 | 201.11.1.3 | Excessive Surface Temperatures | Probe surface | Thermal energy (High temperature) | N | Too high-intensity ultraound energy | Unintended ultrasound energy radiation during diagnosis due to S/W error | Radiating excessive ultrasound energy on the skin surface | Burn | 3 | 2 | 6 | N | D |  |  | 3 | 1 | 3 | Y | N/A | No new and increased risks
[15 empty] | P | Perform acoustic power and index and system acceptance testing to ensure that energy level is lower than defined limit | Acoustic output Report (HRD-01-21)
I
HID-61 | IEC 60601-2-37 | 201.11.1.3.2 | Transmit parameters | Probe surface | [10 empty] | D
P
I
HID-62 | IEC 60601-2-37 | 201.11.1.3.2 | Measurement uncertainty | Probe surface | [10 empty] | D
P
I
HID-63 | IEC 60601-2-37 | 201.12.4.5.1 | Acoustic output limits | Acoustic energy can be exceed when equipment fault | Mechanical energy (Ultrasound energy) | F | Too high-intensity ultraound energy | Unintended ultrasound energy radiation during diagnosis due to S/W error | Radiating excessive ultrasound energy on the skin surface | Burn, cell destruction, cavitation | 3 | 2 | 6 | N | D |  |  | 3 | 1 | 3 | Y | N/A | No new and increased risks
[15 empty] | P | Protective measure option adopted | Acoustic output Report (HRD-01-21)
I
HID-64 | IEC 62304 | 4.2 | Microprocessor | Initalization | Function (Loss or deterioration of function) | N | Firmware upload failed | inoperative the system | Do not use equipment | Diagnostic error | 2 | 5 | 10 | N | D | designed in function to re-analyzed the protective system (SDS-M01, SDS-M02) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HID-65 | IEC 62304 | 4.2 | Microprocessor | Probe Setting | Function (Loss or deterioration of function) | N | Don't connecte the probe | inoperative the system | Do not use equipment | Diagnostic error | 2 | 5 | 10 | N | D | designed in function to re-analyzed the protective system (SDS-M01, SDS-M02) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HID-66 | IEC 62304 | 4.2 | Microprocessor | System option setting | Function (Loss or deterioration of function) | N | Firmware upload failed Miss the connection | inoperative the system | Do not use equipment | Diagnostic error | 2 | 5 | 10 | N | D | designed in function to re-analyzed the protective system (SDS-M01, SDS-M02) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HID-67 | IEC 62304 | 4.2 | Memory device | Patient information | Function (Loss or deterioration of function) | N | Lack of capacity | Cannot save information | Accurate information can not be obtained. | Diagnostic error | 2 | 5 | 10 | N | D | designed in function to optimize the capacity of logical space (SDS-C01~C09) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HID-68 | IEC 62304 | 4.2 | Memory device | Save Cine Loop | Function (Loss or deterioration of function) | N | Lack of capacity | Cannot save information | Accurate information can not be obtained. | Diagnostic error | 2 | 5 | 10 | N | D | designed in function to optimize the capacity of logical space (SDS-C01~C09) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HID-69 | IEC 62304 | 4.2 | Memory device | Save image | Function (Loss or deterioration of function) | N | Lack of capacity | Don't save information | Accurate information can not be obtained. | Diagnostic error | 2 | 5 | 10 | N | D | designed in function to optimize the capacity of logical space (SDS-C01~C09) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HID-70 | IEC 62304 | 4.2 | Sensor | Probe slection | Function (Loss or deterioration of function) | N | Don't connecte the probe | inoperative the system | Do not use equipment | Diagnostic error | 2 | 5 | 10 | N | D | 1. designed in function to re-analyzed the protective system (SDS-M01, SDS-M02) 2. designed in function to check the status of connection (SDS-E01 ~ S04) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HID-71 | IEC 62304 | 4.2 | Energy source | Tx | Function (Loss or deterioration of function) | N | Don't connecte the probe | inoperative the system | Accurate information can not be obtained. | Diagnostic error | 2 | 5 | 10 | N | D | 1. designed in function to re-analyzed the protective system (SDS-M01, SDS-M02) 2. designed in function to check the status of connection (SDS-E01 ~ S04) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HID-72 | IEC 62304 | 4.2 | Energy source | Rx | Function (Loss or deterioration of function) | N | Don't connecte the probe | inoperative the system | Accurate information can not be obtained. | Diagnostic error | 2 | 5 | 10 | N | D | 1. designed in function to re-analyzed the protective system (SDS-M01, SDS-M02) 2. designed in function to check the status of connection (SDS-E01 ~ S04) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HID-73 | IEC 62304 | 4.2 | Energy source | Channel Acq | Function (Loss or deterioration of function) | N | Wrong calculation process | inoperative the system | Accurate information can not be obtained. | Diagnostic error | 2 | 5 | 10 | N | D | designed in function to check to the verification cacluation (SDS-E18~E22) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) )
I
HID-74 | IEC 62304 | 4.2 | Energy source | Frequency | Function (Loss or deterioration of function) | N | Wrong calculation process | inoperative the system | Accurate information can not be obtained. | Diagnostic error | 2 | 5 | 10 | N | D | designed in function to check to the verification cacluation (SDS-E18~E22) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) )
I
HID-75 | IEC 62304 | 4.2 | Safety features | MI | Function (Loss or deterioration of function) | N | Too high MI Value | Outside the Value | Affect to the patient. | Diagnostic error | 2 | 5 | 10 | N | D | designed in function to indicate MI, TI index on the screen (SDS-S01~02) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 2 | 4 | Y | N/A | No new and increased risks
P
I
HID-76 | IEC 62304 | 4.2 | Safety features | TI | Function (Loss or deterioration of function) | N | Too high TI Value | Outside the Value | Affect to the patient. | Diagnostic error | 2 | 5 | 10 | N | D | designed in function to indicate MI, TI index on the screen (SDS-S01~02) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 2 | 4 | Y | N/A | No new and increased risks
P
I
HID-78 | IEC 62304 | 4.2 | Communications | Freeze | Function (Loss or deterioration of function) | N | User doesn't check the manual | Don't operate the Freeze | Do not use equipment | Diagnostic error | 1 | 5 | 5 | N | D | 1. designed in function to re-analyzed the protective system (SDS-M01, SDS-M02) 2. designed in function to check the status of connection (SDS-E01 ~ S04) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 1 | 3 | 3 | Y | N/A | No new and increased risks
P
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 2. Getting started)
HID-79 | IEC 62304 | 4.2 | Communications | Vertical | Function (Loss or deterioration of function) | N | User doesn't check the manual | Don't operate the Vertical | Do not use equipment | Diagnostic error | 1 | 5 | 5 | N | D | designed in function to re-analyzed the protective system (SDS-M01, SDS-M02) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 1 | 3 | 3 | Y | N/A | No new and increased risks
P
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 2. Getting started)
HID-80 | IEC 62304 | 4.2 | Communications | Horizontal | Function (Loss or deterioration of function) | N | User doesn't check the manual | Don't operate the Horizontal | Do not use equipment | Diagnostic error | 1 | 5 | 5 | N | D | designed in function to re-analyzed the protective system (SDS-M01, SDS-M02) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 1 | 3 | 3 | Y | N/A | No new and increased risks
P
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 2. Getting started)
HID-81 | IEC 62304 | 4.2 | Communications | Dynamic Range | Function (Loss or deterioration of function) | N | User doesn't check the manual | Don't operate the Dynamic Range | Do not use equipment | Diagnostic error | 2 | 5 | 10 | N | D | designed in function to re-analyzed the protective system (SDS-M01, SDS-M02) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 2 | 4 | Y | N/A | No new and increased risks
P
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 2. Getting started)
HID-82 | IEC 62304 | 4.2 | Communications | Patient Input | Function (Loss or deterioration of function) | N | User doesn't check the manual | Don't operate the Patint option | Do not use equipment | Diagnostic error | 2 | 5 | 10 | N | D | designed in function to re-analyzed the protective system (SDS-M01, SDS-M02) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) )
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 2. Getting started)
HID-83 | IEC 62304 | 4.2 | Communications | Focus# | Function (Loss or deterioration of function) | N | User doesn't check the manual | Don't operate the Focus and Number | Do not use equipment | Diagnostic error | 2 | 5 | 10 | N | D | designed in function to re-analyzed the protective system (SDS-M01, SDS-M02) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 2. Getting started)
HID-84 | IEC 62304 | 4.2 | Communications | Zoom | Function (Loss or deterioration of function) | N | User doesn't check the manual | Don't operate the Zoom | Do not use equipment | Diagnostic error | 2 | 5 | 10 | N | D | designed in function to re-analyzed the protective system (SDS-M01, SDS-M02) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 2. Getting started)
HID-85 | IEC 62304 | 4.2 | Communications | Mode Selection | Function (Loss or deterioration of function) | N | User doesn't check the manual | Don't operate the Probe selection | Do not use equipment | Diagnostic error | 2 | 5 | 10 | N | D | 1. Function to check logic of the Probe Selection (SDS-I07) 2. Function to check logic of connection between modules. (SDS-C09) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 2. Getting started)
HID-86 | IEC 62304 | 4.2 | External equipment | DICOM | Function (Loss or deterioration of function) | N | User doesn't check the manual | Don't sent to the DICOM information | Sharing information between each other is not possible. | Diagnostic error | 2 | 5 | 10 | N | D | 1. Function to check logic of Lan cable connection. (SDS-I07) 2. Function to check logic of connection between modules. (SDS-C09) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 2. Getting started)
HID-87 | IEC 62304 | 4.2 | External equipment | External Printer | Function (Loss or deterioration of function) | N | User doesn't check the manual | Don't operate external Printer | The printer can not output | Diagnostic error | 2 | 5 | 10 | N | D | 1. Function to check logic of printer cable connection. (SDS-I01) 2. Function to check logic of connection between modules. (SDS-C09) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 2. Getting started)
HID-88 | IEC 62304 | 4.2 | Imaging Processing and Motion (H/W) | Demodulator | Function (Loss or deterioration of function) | N | Firmware upload failed Miss the connection | inoperative the system | Accurate information can not be obtained. | Diagnostic error | 2 | 5 | 10 | N | D | designed in function to re-analyzed the protective system (SDS-M01, SDS-M02) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) )
I
HID-89 | IEC 62304 | 4.2 | Imaging Processing and Motion (H/W) | Clutter Filter & Hilbert | Function (Loss or deterioration of function) | N | Firmware upload failed Miss the connection | inoperative the system | Accurate information can not be obtained. | Diagnostic error | 2 | 5 | 10 | N | D | designed in function to re-analyzed the protective system (SDS-M01, SDS-M02) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) )
I
HID-90 | IEC 62304 | 4.2 | Imaging Processing and Motion (H/W) | ADC | Function (Loss or deterioration of function) | N | Firmware upload failed Miss the connection | inoperative the system | Accurate information can not be obtained. | Diagnostic error | 2 | 5 | 10 | N | D | designed in function to re-analyzed the protective system (SDS-M01, SDS-M02) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) )
I
HID-91 | IEC 62304 | 4.2 | Interface Requirements | Printers | Function (Loss or deterioration of function) | N | Miss the connection | Don't operate the Printer | Do not use the Printer | Diagnostic error | 2 | 5 | 10 | N | D | 1. Function to check logic of printer cable connection. (SDS-I01) 2. Function to check logic of connection between modules. (SDS-C09) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HID-92 | IEC 62304 | 4.2 | Interface Requirements | Monitors | Function (Loss or deterioration of function) | N | Miss the connection | Don't operate the Monitor | Do not use the Monitor | Diagnostic error | 2 | 5 | 10 | N | D | 1. Function to check logic of the monitor connection. (SDS-I02) 2. Function to check logic of connection between modules. (SDS-C09) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HID-93 | IEC 62304 | 4.2 | Interface Requirements | Keyboard | Function (Loss or deterioration of function) | N | Miss the connection | Don't operate the keborad | Do not use the keyboard | Diagnostic error | 2 | 5 | 10 | N | D | 1. Function to check logic of the keyboard connection (SDS-I03) 2. Function to check logic of connection between modules. (SDS-C09) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
| IEC 62304 |  | Interface Requirements | Mouse |  | N | Miss the connection | Don't operate the Mouse | Do not use the mouse |  | 2 |  |  |  | I
HID-94 | IEC 62304 | 4.2 | Interface Requirements | Mouse | Function (Loss or deterioration of function) | N | Miss the connection | Don't operate the Mouse. | Accurate information can not be obtained. | Diagnostic error | 2 | 5 | 10 | N | D | 1. Function to check logic of the trackball connection (SDS-I03) 2. Function to check logic of connection between modules. (SDS-C09) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
[7 empty] | Miss the connection | [7 empty] | I
HID-95 | IEC 62304 | 4.2 | Interface Requirements | Functions for checking DB size & compaction DB | Function (Loss or deterioration of function) | N | Lack of capacity | Don't save information | Accurate information can not be obtained. | Diagnostic error | 2 | 5 | 10 | N | D | designed in function to optimize the capacity of logical space (SDS-C01~C09) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HID-96 | IEC 62304 | 4.2 | Interface Requirements | LAN | Function (Loss or deterioration of function) | N | Miss the connection | Don't operate the LAN | Do not use the LAN | Diagnostic error | 2 | 5 | 10 | N | D | 1. Function to check logic of Lan cable connection. (SDS-I07) 2. Function to check logic of connection between modules. (SDS-C09) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
[7 empty] | Miss the connection | [7 empty] | I
HID-97 | IEC 62304 | 4.2 | Interface Requirements | VGA | Function (Loss or deterioration of function) | N | Miss the connection | Don't operate the VGA | Do not use the VGA | Diagnostic error | 2 | 5 | 10 | N | D | 1. Function to check logic of the VGA connection (SDS-I09) 2. Function to check logic of connection between modules. (SDS-C09) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
[7 empty] | Miss the connection | [7 empty] | I
HID-98 | IEC 62304 | 4.2 | Interface Requirements | TGC | Function (Loss or deterioration of function) | N | Miss the connection | Don't operate the TGC | Do not use the TGC | Diagnostic error | 2 | 5 | 10 | N | D | 1. Function to check logic of the TGC connection (SDS-I10) 2. Function to check logic of connection between modules. (SDS-C09) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
[7 empty] | Miss the connection | [7 empty] | I
HID-99 | IEC 62304 | 4.2 | Timing and memory requirements | Memory Storage | Function (Loss or deterioration of function) | N | Lack of capacity | Don't save information | Accurate information can not be obtained. | Diagnostic error | 2 | 5 | 10 | N | D | designed in function to optimize the capacity of logical space (SDS-C01~C09) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) )
I
HID-100 | IEC 62304 | 4.2 | Timing and memory requirements | Memory Storage | Function (Loss or deterioration of function) | N | Lack of capacity | Don't save information | Accurate information can not be obtained. | Diagnostic error | 2 | 5 | 10 | N | D | designed in function to optimize the capacity of logical space (SDS-C01~C09) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) )
I
HID-101 | IEC 62304 | 4.2 | Timing and memory requirements | Memory Storage | Function (Loss or deterioration of function) | N | Lack of capacity | Don't save information | Accurate information can not be obtained. | Diagnostic error | 2 | 5 | 10 | N | D | designed in function to optimize the capacity of logical space (SDS-C01~C09) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) )
I
HID-102 | IEC 62304 | 4.2 | Imaging Processing and Motion (S/W) | DAS (Digital scan conversion) | Function (Loss or deterioration of function) | N | Firmware upload failed Miss the connection | Wrong output the screen | Accurate information can not be obtained. | Diagnostic error | 2 | 5 | 10 | N | D | designed in function to re-analyzed the protective system (SDS-M01, SDS-M02) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) )
I
HID-103 | IEC 62304 | 4.2 | Imaging Processing and Motion (S/W) | Envelop Detection | Function (Loss or deterioration of function) | N | Firmware upload failed Miss the connection | Wrong calculate the gain | Accurate information can not be obtained. | Diagnostic error | 2 | 5 | 10 | N | D | designed in function to re-analyzed the protective system (SDS-M01, SDS-M02) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) )
I
HID-104 | IEC 62304 | 4.2 | Imaging Processing and Motion (S/W) | Log Compression | Function (Loss or deterioration of function) | N | Firmware upload failed Miss the connection | Wrong calculate the gain | Accurate information can not be obtained. | Diagnostic error | 2 | 5 | 10 | N | D | designed in function to re-analyzed the protective system (SDS-M01, SDS-M02) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) )
I
HID-105 | IEC 62304 | 4.2 | Imaging Processing and Motion (S/W) | Distal Scan Conversion | Function (Loss or deterioration of function) | N | Firmware upload failed Miss the connection | Do not want to screen | Accurate information can not be obtained. | Diagnostic error | 2 | 5 | 10 | N | D | designed in function to re-analyzed the protective system (SDS-M01, SDS-M02) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) )
I
HID-106 | IEC 62304 | 4.2 | Imaging Processing and Motion (S/W) | Base Line | Function (Loss or deterioration of function) | N | Firmware upload failed | It's not a calculation. | Accurate information can not be obtained. | Diagnostic error | 2 | 5 | 10 | N | D | designed in fuction to idetifiy and check the screen interface (SDS-M03) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HID-107 | IEC 62304 | 4.2 | Imaging Processing and Motion (S/W) | Body mark | Function (Loss or deterioration of function) | N | Firmware upload failed Miss the connection | It's not a appearance on the monitor | Don't monitoring | Diagnostic error | 2 | 5 | 10 | N | D | designed in fuction to idetifiy and check the screen interface (SDS-M03) | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HID-108 | IEC 62366 | E2.22 | Is the MEDICAL DEVICE intended to be mobile or portable? | Mobile ME Equipment | Mechanical energy (Moving parts) | N | Lack of Wheel-lock's performance. | Brake failure | Physical impact on person by the equipment | Physical injury | 3 | 2 | 6 | N | D | Using lock-type to all casters | Design V&V(Q5-29-041) | 3 | 1 | 3 | Y | N/A | No new and increased risks
P
I
HID-109 | IEC 62366 | E2.16 | Does the MEDICAL DEVICE have connecting parts or accessories? | Storage of the probes | Use error (Knowledge-based failure) | N | Placing the probe at the inadequate location | Contamination of the Probe | Patients infected by contaminated probes | Skin trouble | 2 | 4 | 8 | N | D | Using probes meeting with related public certificate | Critical component list (No. 277190) | 2 | 2 | 4 | Y | N/A | No new and increased risks
P
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 9. Troubleshooting and Maintenance)
HID-110 | IEC 62366 | E2.12 | Does installation or use of the MEDICAL DEVICE require special training or special skills? | Reading information of the display | Mechanical energy (Moving parts) | N | Inadequate manipulation | Installation failure | Physical impact on person by the equipment | Physical injury | 3 | 1 | 3 | Y | D |  |  | 3 | 1 | 3 | Y | N/A | No new and increased risks
P
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 2. Getting started)
HID-111 | IEC 62366 | E2.14 | Can the USER INTERFACE design features contribute to USE ERROR? | Operating Control panel | Mechanical energy (Ultrasound energy) | N | User presses the wrong button | Diagnostic error | Repetition of diagnosis | Inconvenience during diagnosis | 2 | 2 | 4 | Y | D | designed in consideration of usability | Safety Test Report (No. 277190) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 2. Getting started)
HID-112 | IEC 62366 | E.2.21 | In what way(s) might the MEDICAL DEVICE be deliberately misused? | Cleaning | Use error (Knowledge-based failure) | N | Cleaning method is not specified | Use while in a contaminated status | Patients infected by contaminated substances | Skin trouble | 2 | 4 | 8 | N | D | Using material meeting with related public certificate | Critical component list (No. 277190) | 2 | 2 | 4 | Y | N/A | No new and increased risks
P
[15 empty] | I | Description of cleaning method and warnings in manual | User Manual (HM-01-0 Chapter 9. Troubleshooting and Maintenance)
HID-113 | IEC 62366 | E2.2 | Are measurements taken? | Gripping/holding the probe device | Use error (Knowledge-based failure) | N | Inappropriate gripping and holding probes | Diagnostic error | Repetition of diagnosis | Inconvenience during diagnosis | 2 | 2 | 4 | Y | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 2. Getting started)
HID-114 | IEC 62366 | E2.16 | Does the MEDICAL DEVICE have connecting parts or accessories? | Connecting Probe | Electromagnetic energy (Line voltage, Leakage current) | N | Connecting the probe during no-freezing state | Failure of the PSA PCB | Electric shock due to contact with the PCB directly | Elect. Shock | 3 | 4 | 12 | N | D | Using insulated material meeting with related UL certificate | Critical component list (No. 277190) | 3 | 1 | 3 | Y | N/A | No new and increased risks
[15 empty] | P | Performing a safety test: 1. resistance to moisture 2. dielectric strength 3. mechanical strength | Safety Test Report (No. 277190)
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 10. Safety)
HID-115 | IEC 62366 | E2.16 | Does the MEDICAL DEVICE have connecting parts or accessories? | Detaching Probe | Electromagnetic energy (Line voltage, Leakage current) | N | Detaching the probe during no-freezing state | Failure of the PSA PCB | Electric shock due to contact with the PCB directly | Elect. Shock | 3 | 4 | 12 | N | D | Using insulated material meeting with related UL certificate | Critical component list (No. 277190) | 3 | 1 | 3 | Y | N/A | No new and increased risks
[15 empty] | P | Performing a safety test: 1. resistance to moisture 2. dielectric strength 3. mechanical strength | Safety Test Report (No. 277190)
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 10. Safety)
HID-116 | IEC 62366 | E2.16 | Does the MEDICAL DEVICE have connecting parts or accessories? | Changing Probe | Use error (Knowledge-based failure) | N | Variety of probe | Wrong purpose to choice the probe | Operation failed due to device performance errors | Diagnostic error | 2 | 1 | 2 | Y | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
[15 empty] | I | Information in the manual | Provision of Technical & General manual Training of operator before use
HID-117 | IEC 62366 | E.2.5 | Is the MEDICAL DEVICE susceptible to environmental influences? | Storage of the ME equipment | Electromagnetic energy (Line voltage, Leakage current) | N | Inappropriate storage of the ME equipment | Failure of the system | Electric shock on operators or patients due to the destroyed insulation. | Elect. Shock | 3 | 2 | 6 | N | D | Using material meeting with related public certificate | Critical component list (No. 277190) | 3 | 1 | 3 | Y | N/A | No new and increased risks
P
[15 empty] | I | Information in the manual | User Manual (HM-01-0 Chapter 8. Specification)
HID-118 | IEC 62366 | E2.13 | How will information for safe use be provided? | Information for the identifications of product | Information Hazard (Labelling) | N | inadequate labeling | Wrong attachment during assembly | Diagnosis by the ME equipment with inadequate labels | Diagnostic error | 2 | 3 | 6 | N | D | Control the process progree of assembly with Assembly Guidance | Assembly Guidance (HRD-08-00) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HID-119 | IEC 62366 | E2.13 | How will information for safe use be provided? | Information for the identifications of product | Information Hazard (Labelling) | N | Low adhesive property | not or barely legible | Diagnosis by the ME equipment without labels | Diagnostic error | 2 | 3 | 6 | N | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Performing a safety test: Durability Marking test | Safety Test Report (No. 277190)
I