## Sheet1
Hazard ID | Regulation | Clause | Part/System Description | 4.2 Characteristic/ Intended use | Risk Analysis | [7 empty] | 5. Risk Evaluation |  | 6. Risk control
[5 empty] | 4.3 Hazard | 4.4 Estimate risk(s) for hazardous situation | [8 empty] | 6.2 Option Analysis |  | 6.3 Implementation of risk control measures | 6.4. Residual risk evaluation |  |  |  | 6.5 Benefit | 6.6 Other Hazards
[6 empty] | Normal/ Fault Condition (N/F)? | Initiating Events and Circumstances/Causes (Potential Cause of Hazard/ Failure) (원인)) | Foreseeable sequences of Events | Hazardous situation | Harm | Initial Risk Estimation
[11 empty] | Sev. | Prob. | Risk | Acceptable Risk?(Y/N) | Risk control Measurment (Description of Mitigation or Criteria) D: Inherent safety by design P: Protective measures in themedical device itself or in the manufacturing process I: Information for safety |  | Verification/ Validation | Sev. | Prob. | Residual Risk | Accptable Risk?(Y/N)
HZ-01 | IEC 62304 | 4.2 | Microprocessor | Initalization | Function (Loss or deterioration of function) | N | Firmware upload failed | inoperative the system | Do not use equipment | Diagnostic error | 2 | 3 | 6 | N | D | Protective system | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HZ-02 | IEC 62304 | 4.2 | Microprocessor | Probe Setting | Function (Loss or deterioration of function) | N | Don't connecte the probe | inoperative the system | Do not use equipment | Diagnostic error | 2 | 3 | 6 | N | D | Protective system | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HZ-03 | IEC 62304 | 4.2 | Microprocessor | System option setting | Function (Loss or deterioration of function) | N | Firmware upload failed Miss the connection | inoperative the system | Do not use equipment | Diagnostic error | 2 | 3 | 6 | N | D | Designing software with reanalyze-function | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HZ-04 | IEC 62304 | 4.2 | Memory device | Patient information | Function (Loss or deterioration of function) | N | Lack of capacity | Cannot save information | Accurate information can not be obtained. | Diagnostic error | 2 | 1 | 2 | Y | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HZ-05 | IEC 62304 | 4.2 | Memory device | Save Cine Loop | Function (Loss or deterioration of function) | N | Lack of capacity | Cannot save information | Accurate information can not be obtained. | Diagnostic error | 2 | 1 | 2 | Y | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HZ-06 | IEC 62304 | 4.2 | Memory device | Save image | Function (Loss or deterioration of function) | N | Lack of capacity | Don't save information | Accurate information can not be obtained. | Diagnostic error | 2 | 1 | 2 | Y | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HZ-07 | IEC 62304 | 4.2 | Sensor | Probe slection | Function (Loss or deterioration of function) | N | Don't connecte the probe | inoperative the system | Do not use equipment | Diagnostic error | 2 | 1 | 2 | Y | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HZ-08 | IEC 62304 | 4.2 | Energy source | Tx | Function (Loss or deterioration of function) | N | Don't connecte the probe | inoperative the system | Accurate information can not be obtained. | Diagnostic error | 2 | 3 | 6 | N | D | Designing software with reanalyze-function | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HZ-09 | IEC 62304 | 4.2 | Energy source | Rx | Function (Loss or deterioration of function) | N | Don't connecte the probe | inoperative the system | Accurate information can not be obtained. | Diagnostic error | 2 | 3 | 6 | N | D | Designing software with reanalyze-function | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HZ-10 | IEC 62304 | 4.2 | Energy source | Channel Acq | Function (Loss or deterioration of function) | N | Wrong calculation process | inoperative the system | Accurate information can not be obtained. | Diagnostic error | 2 | 3 | 6 | N | D | Check to the verification cacluation | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) )
I
HZ-11 | IEC 62304 | 4.2 | Energy source | Frequency | Function (Loss or deterioration of function) | N | Wrong calculation process | inoperative the system | Accurate information can not be obtained. | Diagnostic error | 2 | 3 | 6 | N | D | Check to the verification cacluation | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | ```
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) )
I
HZ-12 | IEC 62304 | 4.2 | Safety features | MI | Function (Loss or deterioration of function) | N | Too high MI Value | Outside the Value | Affect to the patient. | Diagnostic error | 2 | 2 | 4 | Y | D |  |  | 2 | 2 | 4 | Y | N/A | No new and increased risks
P
I
HZ-13 | IEC 62304 | 4.2 | Safety features | TI | Function (Loss or deterioration of function) | N | Too high TI Value | Outside the Value | Affect to the patient. | Diagnostic error | 2 | 2 | 4 | Y | D |  |  | 2 | 2 | 4 | Y | N/A | No new and increased risks
P
I
HZ-14 | IEC 62304 | 4.2 | Communications | Power On/Off | Function (Loss or deterioration of function) | N | User is not checking manual | Don't off the system | Do not use equipment | Diagnostic error | 2 | 3 | 6 | N | D | The protective earth terminals implementation Performin a safety test | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HZ-15 | IEC 62304 | 4.2 | Communications | Freeze | Function (Loss or deterioration of function) | N | User is not checking manual | Don't operate the Freeze | Do not use equipment | Diagnostic error | 1 | 3 | 3 | Y | D |  |  | 1 | 3 | 3 | Y | N/A | No new and increased risks
P
I
HZ-16 | IEC 62304 | 4.2 | Communications | Vertical | Function (Loss or deterioration of function) | N | User is not checking manual | Don't operate the Vertical | Do not use equipment | Diagnostic error | 1 | 3 | 3 | Y | D |  |  | 1 | 3 | 3 | Y | N/A | No new and increased risks
P
I
HZ-17 | IEC 62304 | 4.2 | Communications | Horizontal | Function (Loss or deterioration of function) | N | User is not checking manual | Don't operate the Horizontal | Do not use equipment | Diagnostic error | 1 | 3 | 3 | Y | D |  |  | 1 | 3 | 3 | Y | N/A | No new and increased risks
P
I
HZ-18 | IEC 62304 | 4.2 | Communications | Dynamic Range | Function (Loss or deterioration of function) | N | User is not checking manual | Don't operate the Dynamic Range | Do not use equipment | Diagnostic error | 2 | 2 | 4 | Y | D |  |  | 2 | 2 | 4 | Y | N/A | No new and increased risks
P
I
HZ-19 | IEC 62304 | 4.2 | Communications | Patient Input | Function (Loss or deterioration of function) | N | User is not checking manual | Don't operate the Patint option | Do not use equipment | Diagnostic error | 2 | 3 | 6 | N | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) )
I
HZ-20 | IEC 62304 | 4.2 | Communications | Focus# | Function (Loss or deterioration of function) | N | User is not checking manual | Don't operate the Focus and Number | Do not use equipment | Diagnostic error | 2 | 1 | 2 | Y | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HZ-21 | IEC 62304 | 4.2 | Communications | Zoom | Function (Loss or deterioration of function) | N | User is not checking manual | Don't operate the Zoom | Do not use equipment | Diagnostic error | 2 | 1 | 2 | Y | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HZ-22 | IEC 62304 | 4.2 | Communications | Mode Selection | Function (Loss or deterioration of function) | N | User is not checking manual | Don't operate the Probe selection | Do not use equipment | Diagnostic error | 2 | 3 | 6 | N | D | Function to check the Probe Selection Function to check the Connection | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HZ-23 | IEC 62304 | 4.2 | External equipment | DICOM | Function (Loss or deterioration of function) | N | User is not checking manual | Don't sent to the DICOM information | Sharing information between each other is not possible. | Diagnostic error | 2 | 3 | 6 | N | D | Function to check the Lan Cable. Function to check connection between system. | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HZ-24 | IEC 62304 | 4.2 | External equipment | External Printer | Function (Loss or deterioration of function) | N | User is not checking manual | Don't operate external Printer | The printer can not output | Diagnostic error | 2 | 1 | 2 | Y | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HZ-25 | IEC 62304 | 4.2 | Imaging Processing and Motion (H/W) | Demodulator | Function (Loss or deterioration of function) | N | Firmware upload failed Miss the connection | inoperative the system | Accurate information can not be obtained. | Diagnostic error | 2 | 3 | 6 | N | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) )
I
HZ-26 | IEC 62304 | 4.2 | Imaging Processing and Motion (H/W) | Clutter Filter & Hilbert | Function (Loss or deterioration of function) | N | Firmware upload failed Miss the connection | inoperative the system | Accurate information can not be obtained. | Diagnostic error | 2 | 3 | 6 | N | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) )
I
HZ-27 | IEC 62304 | 4.2 | Imaging Processing and Motion (H/W) | ADC | Function (Loss or deterioration of function) | N | Firmware upload failed Miss the connection | inoperative the system | Accurate information can not be obtained. | Diagnostic error | 2 | 3 | 6 | N | D | function to check the verification cacluation | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) )
I
HZ-28 | IEC 62304 | 4.2 | Interface Requirements | Printers | Function (Loss or deterioration of function) | N | Miss the connection | Don't operate the Printer | Do not use the Printer | Diagnostic error | 2 | 1 | 2 | Y | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HZ-29 | IEC 62304 | 4.2 | Interface Requirements | Monitors | Function (Loss or deterioration of function) | N | Miss the connection | Don't operate the Monitor | Do not use the Monitor | Diagnostic error | 2 | 1 | 2 | Y | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HZ-30 | IEC 62304 | 4.2 | Interface Requirements | Keyboard | Function (Loss or deterioration of function) | N | Miss the connection | Don't operate the keborad | Do not use the keyboard | Diagnostic error | 2 | 3 | 6 | N | D | Well put power protection terminal Connector checks | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
| IEC 62304 |  | Interface Requirements | Mouse |  | N | Miss the connection | Don't operate the Mouse | Do not use the mouse |  | 2 | 1 |  | Y | I
HZ-31 | IEC 62304 | 4.2 | Interface Requirements | Mouse | Function (Loss or deterioration of function) | N | Miss the connection | Don't operate the Mouse. | Accurate information can not be obtained. | Diagnostic error | 2 | 3 | 6 | N | D | Well put power protection terminal Connector checks | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
[7 empty] | Miss the connection | [7 empty] | I
HZ-32 | IEC 62304 | 4.2 | Interface Requirements | Functions for checking DB size & compaction DB | Function (Loss or deterioration of function) | N | Lack of capacity | Don't save information | Accurate information can not be obtained. | Diagnostic error | 2 | 1 | 2 | Y | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HZ-33 | IEC 62304 | 4.2 | Interface Requirements | USB | Function (Loss or deterioration of function) | N | User is not checking manual | Don't operate the USB | Do not use the USB | Diagnostic error | 2 | 3 | 6 | N | D | Well put power protection terminal Connector checks | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HZ-34 | IEC 62304 | 4.2 | Interface Requirements | LAN | Function (Loss or deterioration of function) | N | Miss the connection | Don't operate the LAN | Do not use the LAN | Diagnostic error | 2 | 1 | 2 | Y | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
[7 empty] | Miss the connection | [7 empty] | I
HZ-35 | IEC 62304 | 4.2 | Interface Requirements | VGA | Function (Loss or deterioration of function) | N | Miss the connection | Don't operate the VGA | Do not use the VGA | Diagnostic error | 2 | 1 | 2 | Y | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
[7 empty] | Miss the connection | [7 empty] | I
HZ-36 | IEC 62304 | 4.2 | Interface Requirements | CP (Control Panel) | Function (Loss or deterioration of function) | N | Miss the connection | Don't operate the CP | Do not use the CP | Diagnostic error | 2 | 3 | 6 | N | D | Well put power protection terminal Function to check onnector | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) ) | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
[7 empty] | Miss the connection | [7 empty] | I
HZ-37 | IEC 62304 | 4.2 | Interface Requirements | TGC | Function (Loss or deterioration of function) | N | Miss the connection | Don't operate the TGC | Do not use the TGC | Diagnostic error | 2 | 1 | 2 | Y | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
[7 empty] | Miss the connection | [7 empty] | I
HZ-38 | IEC 62304 | 4.2 | Timing and memory requirements | Memory Storage | Function (Loss or deterioration of function) | N | Lack of capacity | Don't save information | Accurate information can not be obtained. | Diagnostic error | 2 | 1 | 2 | Y | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HZ-39 | IEC 62304 | 4.2 | Timing and memory requirements | Memory Storage | Function (Loss or deterioration of function) | N | Lack of capacity | Don't save information | Accurate information can not be obtained. | Diagnostic error | 2 | 1 | 2 | Y | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HZ-40 | IEC 62304 | 4.2 | Timing and memory requirements | Memory Storage | Function (Loss or deterioration of function) | N | Lack of capacity | Don't save information | Accurate information can not be obtained. | Diagnostic error | 2 | 1 | 2 | Y | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HZ-41 | IEC 62304 | 4.2 | Imaging Processing and Motion (S/W) | DAS (Digital scan conversion) | Function (Loss or deterioration of function) | N | Firmware upload failed Miss the connection | Wrong output the screen | Accurate information can not be obtained. | Diagnostic error | 2 | 3 | 6 | N | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) )
I
HZ-42 | IEC 62304 | 4.2 | Imaging Processing and Motion (S/W) | Envelop Detection | Function (Loss or deterioration of function) | N | Firmware upload failed Miss the connection | Wrong calculate the gain | Accurate information can not be obtained. | Diagnostic error | 2 | 3 | 6 | N | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) )
I
HZ-43 | IEC 62304 | 4.2 | Imaging Processing and Motion (S/W) | Log Compression | Function (Loss or deterioration of function) | N | Firmware upload failed Miss the connection | Wrong calculate the gain | Accurate information can not be obtained. | Diagnostic error | 2 | 3 | 6 | N | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) )
I
HZ-44 | IEC 62304 | 4.2 | Imaging Processing and Motion (S/W) | Distal Scan Conversion | Function (Loss or deterioration of function) | N | Firmware upload failed Miss the connection | Do not want to screen | Accurate information can not be obtained. | Diagnostic error | 2 | 3 | 6 | N | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
[15 empty] | P | Warning to the user by the Error Message | Software Validation report (Q5-29-028(4), 6. Software Design Specification(SDS) )
I
HZ-45 | IEC 62304 | 4.2 | Imaging Processing and Motion (S/W) | Base Line | Function (Loss or deterioration of function) | N | Firmware upload failed | It's not a calculation. | Accurate information can not be obtained. | Diagnostic error | 2 | 3 | 6 | Y | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I
HZ-46 | IEC 62304 | 4.2 | Imaging Processing and Motion (S/W) | Body mark | Function (Loss or deterioration of function) | N | Firmware upload failed Miss the connection | It's not a appearance on the monitor | Don't monitoring | Diagnostic error | 2 | 1 | 2 | Y | D |  |  | 2 | 1 | 2 | Y | N/A | No new and increased risks
P
I